Figure 2. An example experiment structure.
OpenSesame EyeLink tutorial
Daniel Schreij1, Sebastiaan Mathôt1,2, and Lotje van der Linden2
1 VU University Amsterdam, Dept. of Cognitive Psychology
2 Aix-Marseille Université, Laboratoire de Psychologie Cognitive
Last updated 11/28/12
http://www.cogsci.nl/opensesame
# 1 Communicating with the EyeLink through the OpenSesame GUI
Creating a basic eye-tracking experiment using the drag-and-drop interface of OpenSesame (Mathôt, Schreij, & Theeuwes, 2012) is very easy. After installing the OpenSesame plug-ins <http://osdoc.cogsci.nl/plug-ins/eyelink-plug-ins> for communicating with the SR Research EyeLink, you will see that the some extra icons have appeared the item toolbar (Figure 1).

These icons represent the various options you have for communicating with the EyeLink device. Of course, it is also possible to communicate with the EyeLink directly through the scripting interface, but this is described in a later section.
Figure 1. The Eyelink plug-in icons in the item toolbar
Basically, an experiment using the EyeLink consists of the following phases.
## 1.1 Connecting to the Eyelink (once per experiment)
### eyelink_calibrate

Before you can start measuring the eyes, you first have to establish a connection with the EyeLink and then need to calibrate it, so it can associate the position of the measured eyes with coordinates on the screen. This is all done with the eyelink_calibrate module, of which the icon is depicted above. This module automatically connects to the EyeLink, which you typically only need to do once per experiment, so this module is usually also inserted only once at the top of the experiment tree (at the very beginning of the experiment). Once the connection is established you get into calibration mode, in which you will be presented with the (probably familiar) EyeLink calibration screen. Here you have the following options:
- C – Start the calibration
- V – Validate the calibration you just performed (hence always after having pressed C first)
- Enter / Return – Switch between eye camera view and calibration mode
- A – Automatically determine the pupil detection threshold
- Q – Exit the calibration mode and start or continue the experiment.
If measurements suddenly become less accurate, you always have the possibility to perform a recalibration during drift correction (see 1.2). The eyelink_calibrate module further offers the following options:
- Tracker attached (yes|no): indicates whether the EyeLink should be really approached, or if this is just a ‘dummy run’ of the experiments, during which no real eye measurements should be performed yet. This is useful for testing or debugging an experiment in which you have already implemented all communications with the EyeLink, but would like to perform a test run on a machine without an EyeLink attached.
- Saccade Velocity Threshold (default: 35): indicates how fast an eye movement needs to be, before it is registered as a saccade and indexed in the saccade report by the EyeLink data viewer. It can usually be left at its default value. This value can also be adjusted in the EyeLink data viewer after the data has been collected.
- Saccade Acceleration Threshold (default: 9500): indicates how large the acceleration of an eye movement needs to be before it is registered as a saccade. This can usually be left at its default value.
## 1.2 Drift correction (before every block or trial)
### eyelink_drift_correct
There is always a small drift, due to factors such as movement of the head and slippage of the head gear (in case of a head-mounted tracker). This can slightly offset the calibration. Luckily, you can check whether this is the case wherever you have placed the eyelink_drift_correct module in your experiment. When the EyeLink is in drift correction mode it pauses the experiment until the participants fixates on the center dot on the screen and optionally presses space to continue. While the EyeLink is paused in drift correction mode, it is also possible to recalibrate it again by pressing Q. This brings you back into calibration mode with its options as described in section 1.1.
There is an important difference between the EyeLink 1000 and previous versions of the device (EyeLink 1 and 2). For the Eyelink 1 and 2, drift correction includes a slight 1-point recalibration. This is necessary, because slippage of the helmet causes a systematic drift that needs to be compensated for. In contrast, the EyeLink 1000 only checks whether fixation is stable and accurate, but does not adjust the calibration. The rationale behind this is that, because the Eyelink 1000 is not head-mounted, there is no systematic drift that accumulates throughout the experiment, although there is random drift that varies from trial to trial. In other words, for the EyeLink 10000 drift correction should be thought of as a verification process (and thus optional, but recommended), whereas for the EyeLink 1 and 2 drift correction is an active recalibration (and thus crucial).
There is no real convention as to how often you have to perform drift correction. Most commonly it is done at the start of every trial, but (for the EyeLink 1 and 2) some have claimed this has introduced noise in their data (for instance in some motivational experiments) and once at the start of every block is better. You will just have to be the judge of what works best for your experiment.
## 1.3 Measuring eye data (each trial)
### eyelink_start_recording

Place this module where you want to start recording with the eye-tracker, usually at the beginning of a trial. This module has one customizable option:
Log message (default: start_trial): when recording is started, OpenSesame will send the message you specify here to the EyeLink log file, to mark the beginning of the trial.

### eyelink_wait

This module pauses the experiment, until a certain event occurs. The type of event can be selected from the module’s drop-down list, which comprises:
- Saccade start
- Saccade end
- Fixation start
- Fixation end
- Blink start
- Blink end

### eyelink_stop_recording

This module terminates the EyeLink recording process for the current trial and sends the message specified under Log_message to the EyeLink log file, to mark the end of the trial.
## 1.4 Sending custom messages to the EyeLink
### eyelink_log

You also have the ability to send the EyeLink custom messages, which will be stored in the log file. This is especially useful if you want to specify periods of interest in the data file afterwards. For instance, if you start recording at the beginning of a trial, but are only interested in the eye movements after the appearance of a cue that appears somewhere during a trial, you could send the message “SHOWING CUE” to the EyeLink at the moment the cue is presented. The EyeLink data viewer then enables you to filter eye movements from the moment that the “SHOWING CUE” message was sent till the end of the trial. This possibility gives you more freedom as you can just measure eye movements of the entire trial and afterwards make a selection of the collected data that you actually want to use for further analysis.
## 1.5 Example experiment structure
The structure of an experiment that makes use of the EyeLink could look something like Figure 2.

The very first thing that happens is establishing a connection with the EyeLink and calibrating it (eyelink_calibrate). This connection will stay active for the remainder of the experiment and hence only needs to be established once at the beginning. After the welcome/instruction screen has been shown, the experiment commences with the first block and trial. Each trial begins with drift correction (eyelink_drift_correct), during which the experimenter also has the option to recalibrate the EyeLink again, if necessary. After drift correction has been successfully completed (e.g. when the participant fixated on the screen center and pressed space to continue), the command is sent to EyeLink to start measuring eye movements (eyelink_start_recording). Immediately thereafter the stimulus display is shown. After the participant has given his response, a feedback display is shown and any variables that still needed to be logged are sent to the EyeLink (eyelink_log). Here you can also log all your relevant (independent) variables, like the properties of your target, simply by using the '[variable name]' notation (for an alternative, see 3.1). In this case your log messages would look something like this:
var target_color [col]
var target_x_coordinate [x_pos]
 Finally, the command is sent to the EyeLink to stop recording and finish its log for the current trial. The connection to the EyeLink is automatically closed when your experiment finishes, so no need to worry about that!
# 2 Communicating with the EyeLink through the scripting interface
You can also access the EyeLink using Python in an inline_script item. The EyeLink class is a property of the experiment class, which (in an inline_script item) you can access like this:
self.experiment.eyelink # For all versions of OpenSesame
exp.eyelink # Shortcut for OpenSesame 0.26 and later
Most functions of this class perform the same actions as the modules described in the previous section. The eyelink_calibrate module for instance has its equivalent in the self.experiment.eyelink.calibrate() function. However, there are many functions which have no corresponding modules in the GUI and the scripting interface hence offers extra functionality.
Another convenient function is the equivalent of the eyelink_log module:
self.experiment.eyelink.log("Message for the EDF")
This function is especially useful if multiple events of interest are handled by the same inline_script item, such as displaying a fixation dot and, after a certain duration, a stimulus. In this case you might want your output file to contain a message for each of those events, which can be achieved simply by calling the above function, with an appropriate message, immediately after you drew something to the display. For example:
my_canvas1.show()
self.experiment.eyelink.log("display fixation")
self.sleep(self.get(“SOA”))
my_canvas2.show()
self.experiment.eyelink.log("display stimulus")
Some other examples:
time, start_pos = self.experiment.eyelink.wait_for_saccade_start()
time, start_pos, end_pos = self.experiment.eyelink.wait_for_saccade_end()
x, y = self.experiment.eyelink.sample()
The full list of available functions can be found online <http://osdoc.cogsci.nl/plug-ins/eyelink-plug-ins>.
# 3 Advanced communications with the EyeLink
The EyeLink logging function provides advanced options for registering experimental variables, interest areas and other things. You can do this both with the self.experiment.eyelink.log function and with the eyelink_log module. Whenever you want the EyeLink to perform a special operation with your message you have to start your message with the !V command, followed by the action you want it to take. Below are some cases that illustrate the available options. It is important to note that the !V commands only work when they are issued while the EyeLink is recording (and hence between eyelink_start_recording and eyelink_stop_recording modules). It is possible to issue certain commands outside the recording phase, but this falls outside the scope of this tutorial. There is a lot which can be done with these commands and a full list can be found in the EyeLink Data viewer help file, under the chapter Protocol for EyeLink Data to Viewer Integration. We recommend reading this chapter if you really want to use the full potential of the command directives which the EyeLink interface has to offer.
## 3.1 Storing trial variables
You might want to keep a record of the variables and values that you used during a trial. This can be done through the command:
!V TRIAL_VAR <variable name> <variable value>
For instance, if you want to register that the target color in that specific trial was red, you can issue the command:
self.experiment.eyelink.log(“!V TRIAL_VAR target red”)
Make sure that you specify unique variable names, otherwise the last specified variable might overwrite an earlier registered variable with the same name.
## 3.2 Designating interest areas
You also have the possibility to specify the interest areas in advance. Interest areas are specific regions on the screen of which you are interested if people have fixated inside them (and how frequent) or made a saccade in their direction. For instance, if your experiment contains a spatial cue and a target that appears at another position, you might want to specify interest areas around these elements to determine later if the eyes actually fixated on the cue or target. With an Interest Area Report created in the EyeLink data viewer you then get an ordered list of which interest areas were visited by the eyes along with all the other related information (fixation duration, saccade start point, etc.)
You can specify an interest area with the command:
!V IAREA <shape> <index> <left x> <top y> <right x> <bottom y> [label]
You can give an interest area two shapes, for which the syntax of the command differs slightly. The first possibility is a rectangle shape, which is for instance referenced like this:
self.experiment.eyelink.log(“!V IAREA RECTANGLE 1 10 5 20 15 cue”)
The above command will log a rectangular interest area stretching from x coordinates 10 to 20 and y coordinates 5 to 15. You also need to specify an index for each interest area that you make (in this case 1). Simply start with an index of 1 and increment this number for each new interest area that you create. The label string at the end is optional, but we recommend that you always specify this.
If we would like to create a second interest area with the shape of an ellipse, the command to issue would look something like this:
self.experiment.eyelink.log(“!V IAREA ELLIPSE 2 300 200 400 300 target”)
It is also possible to specify a custom shape for an interest area. This is has a slightly different syntax and is designated by specifying FREEHAND at the shape variable.
!V IAREA FREEHAND <id> <x1, y1> <x2, y2 > ... <xn, yn> [label]
The label string parameter once again is optional. <xi, yi> refers to the coordinates of a point. The x, y coordinates of each point are separated by a comma.
# 4 References
Mathôt, S., Schreij, D., & Theeuwes, J. (2012). OpenSesame: An open-source, graphical experiment builder for the social sciences. Behavior Research Methods, 44, 314-324.